# Problem 3 — Couette Viscometer

Question

A Couette viscometer places fluid between concentric cylinders; one cylinder rotates and the torque is measured. The inner cylinder rotates at ω=120 rad/s, and the outer cylinder is stationary. The viscometer has the following dimensions and measured torque:

- Inner radius Ri=0.025 m

- Outer radius Ro=0.0255 m, so the gap is δ=Ro−Ri=5.0×10−4 m

- Clearance at the bottom of the inner cylinder h=0.0255 m

- Height of liquid in the annular space H=0.080 m

- Measured torque T=0.092 N⋅m

Find the viscosity μ.

Couette viscometer: inner cylinder rotating at ω inside a stationary outer cylinder

Given: Ri, Ro, δ, h, H, ω and T as listed above; the outer cylinder is stationary.

Additional assumptions: the fluid is Newtonian, so a single viscosity μ describes it; the flow is steady and laminar; the gap is thin compared with the radius (δ/Ri=0.02), so the velocity varies linearly across the side gap and across the bottom clearance; end effects at the free surface are negligible.

## 1 Select the working equation

The inner cylinder has a finite clearance h below it. The measured torque therefore includes shear on the cylinder’s side wall and shear on its bottom face. The slides give the viscosity for this case, with R1=Ri and R2=Ro. This formula appears on the slides without an equation number

μ=2Ro−RihTπRi2 ω4HhRo+Ri2Ro−Ri

The first term in the brackets, 4HhRo, comes from the side wall. The second term, Ri2Ro−Ri, comes from the bottom face.

## 2 Evaluate the numerator

Substitute the gap, clearance and torque

2Ro−RihT=25.0×10−4 m0.0255 m0.092 N⋅m

2Ro−RihT=2.346×10−6 N⋅m3

## 3 Evaluate the bracketed geometric term

Calculate the side-wall term

4HhRo=40.080 m0.0255 m0.0255 m=2.0808×10−4 m3

Calculate the bottom-face term

Ri2Ro−Ri=0.025 m25.0×10−4 m=6.25×10−4 m25.0×10−4 m=3.125×10−7 m3

Add the two terms

4HhRo+Ri2Ro−Ri=2.0808×10−4 m3+0.003125×10−4 m3=2.0839×10−4 m3

## 4 Evaluate the denominator

Calculate the prefactor

πRi2 ω=π6.25×10−4 m2120 s−1=0.23562 m2/s

The radian is dimensionless, so rad/s is written as s−1. Multiply by the bracketed term from Section 3

πRi2 ω4HhRo+Ri2Ro−Ri=0.23562 m2s2.0839×10−4 m3=4.9101×10−5 m5s

## 5 Calculate the viscosity

Divide the numerator by the denominator

μ=2.346×10−6 N⋅m34.9101×10−5 m5/s=0.04778 N⋅sm2

Convert the units, using 1 N/m2=1 Pa and 1 cP=10−3 Pa⋅s

μ=0.0478 Pa⋅s=47.8 cP

The fluid viscosity is about 0.048 Pa·s, roughly 50 times the viscosity of water at room temperature.

## 6 Check the result

Units check

m⋅m⋅N⋅mm2⋅s−1⋅m3=N⋅m3m5/s=N⋅sm2=Pa⋅s

Fraction of the torque carried by the bottom face

Ri2Ro−Ri4HhRo+Ri2Ro−Ri=3.125×10−7 m32.0839×10−4 m3=0.0015

The bottom face carries only about 0.15% of the torque, because the bottom clearance (h=0.0255 m) is 51 times the side gap (δ=5.0×10−4 m), so the bottom shear rate is small.

For comparison, the ideal case on the slides ignores the space below the rotating cylinder. This formula also appears on the slides without an equation number

μ=TRo2−Ri24πωRi2Ro2H

Substitute the given values

μ=0.092 N⋅m0.0255 m2−0.025 m24π120 s−10.025 m20.0255 m20.080 m=2.323×10−6 N⋅m34.9028×10−5 m5/s=0.0474 Pa⋅s

The two results agree within about 1%. That is consistent with the small bottom-face contribution and the thin-gap approximation in the finite-clearance formula.

Check against the source solution: the source’s final value is incorrect. In the side-wall term 4HhRo, it substitutes the gap δ=5.0×10−4 m in place of the clearance h=0.0255 m. It also reports a denominator of 6.135×10−5, which does not follow from its own substitution; that substitution gives 1.035×10−6 m5/s. With h=0.0255 m substituted correctly, as in Sections 3–5, the viscosity is 0.0478 Pa⋅s, not 0.0382 Pa⋅s.

Model note: the given clearance h=0.0255 m equals the outer radius Ro, which is unusually large for a viscometer bottom gap and may be a typographical carry-over. It is used here as given. Because the bottom face carries so little of the torque, even a much smaller clearance would change the result only modestly. For example, h=5.0×10−4 m would give μ=0.0444 Pa⋅s.